WUALT

Rule-Based Physiological
Distress Detection Engine

Complete Technical Documentation

April 2026 · v1.0

# 1. Executive Summary

The WUALT Distress Detection Engine is a standalone, stateful, rule-based system that evaluates physiological signals from a wearable smart ring to detect stress and distress in real time.

| Property | Value |
| --- | --- |
| Architecture | Rule-based, deterministic, no ML |
| Latency | < 100 ms per evaluation (O(1) complexity) |
| Input | Preprocessed pipeline output (4 signals + metadata) |
| Output | State (normal/stress/distress) + confidence + alert |
| Signals | Heart Rate, SpO2, HRV (RMSSD), Skin Temperature |
| Statefulness | Persistence tracking, baseline warmup, alert rotation |
| Dependencies | Python stdlib only — zero external packages |
| File | distress_engine.py (991 lines, fully self-contained) |

The engine sits downstream of the preprocessing pipeline (preprocessing_pipeline.py), which handles raw sensor ingestion, EMA smoothing, Signal Quality Index (SQI) computation, clinical flag generation, and z-score baseline tracking. The distress engine consumes that processed output and makes a detection decision.

# 2. Why Rule-Based, Not Machine Learning

This is a deliberate architectural choice, not a limitation.

| Criterion | Rule-Based | ML-Based |
| --- | --- | --- |
| Interpretability | Every decision traceable to a specific threshold | Black-box; SHAP/LIME are approximations |
| Auditability | Clinicians can read and verify thresholds against literature | Requires statistical validation studies |
| Regulatory path | Rules map directly to clinical guidelines (BTS/WHO) | Requires drift monitoring, retraining infrastructure |
| Data requirements | Works immediately with zero training data | Requires large, labelled, diverse real-world datasets |
| Failure modes | Predictable — change one number to fix | Unpredictable on unseen distributions |
| Edge cases | Explicitly handled (exercise, COPD, pregnancy) | Must be represented in training data |
| Latency | O(1) — constant time | Depends on model size |

For a safety-critical wearable device in early stage, rule-based detection provides a reliable, auditable, and immediately deployable baseline. ML can be layered on top later, but the rule-based engine remains as the safety net.

# 3. System Architecture

The system has three layers: the smart ring hardware (sensors), the preprocessing pipeline (data cleaning, quality scoring, baseline tracking), and the distress detection engine (rule evaluation, state management, alert generation).

Critical separation:

The distress engine has ZERO imports from the preprocessing pipeline. It consumes a well-defined dictionary contract. This means either component can be replaced independently, the engine can be tested with synthetic data, and different hardware backends can feed the same engine.

# 4. Input Contract

The engine's evaluate() method accepts one argument — a dictionary with four top-level keys:

| Key | Type | Description |
| --- | --- | --- |
| sample | dict | Processed sensor frame: HR, SpO2, HRV, temp, SQI, clinical_flags, accepted |
| zscores | dict | Per-signal z-scores against personal baseline |
| baseline_ready | bool | True after sufficient warmup frames (~30s) |
| window | dict | Rolling 30-second window statistics (mean, var, min, max) |

Key fields the engine uses:

sample.hr, sample.spo2, sample.dyn_acc_mag, sample.temp (vital signs); sample.accepted (gate check); sample.sqi.overall (quality gate); sample.clinical_flags (advisory integration); zscores.hr/.hrv_rmssd/.spo2/.temp (personalized deviation); baseline_ready (cold-start handling); window.window_n (confidence scaling).

# 5. Output Contract

| Field | Type | Description |
| --- | --- | --- |
| state | "normal" ¦ "stress" ¦ "distress" | Current detection state |
| confidence | float 0.0–1.0 | Engine confidence in the detection |
| contributing_signals | list of str | Which signals triggered (e.g., ["hr", "spo2"]) |
| alert.title | str | Short heading (e.g., "Heart rate is up") |
| alert.message | str | Full user-facing message — calm, no jargon |
| alert.severity | "low" ¦ "medium" ¦ "high" | Alert priority level |
| debug.flags | dict of bool | Per-signal flag status |
| debug.scores | dict of float | Per-signal severity scores (0–1) |
| debug.persistence_s | float | Seconds this condition has persisted |
| debug.motion_state | str | "still" / "walking" / "active" / "exercise" |
| debug.flag_count | int | Number of signals currently flagged |
| debug.weighted_score | float | Combined weighted confidence score |
| debug.source | str | "zscore" / "absolute" / "clinical_flag" |

# 6. The Seven-Stage Evaluation Pipeline

Each frame passes through seven sequential stages:

## 6.1 Stage 1: Quality Gate

Two checks prevent decisions on bad data. (1) Frame acceptance: if sample.accepted == False (finger off, sensor failure, charging), return normal with confidence 0.0 and reset persistence. (2) SQI threshold: if sqi.overall < 0.5, data quality too low — same behavior.

Why reset persistence on gate failure: if a user removes their ring for 30 seconds during a stress episode, the persistence counter must not keep ticking. When they put it back on, the condition must re-establish from zero.

## 6.2 Stage 2: Motion Classification

Determines whether the user is exercising so elevated HR can be suppressed as expected behavior.

| Motion State | dyn_acc_mag Range | Interpretation |
| --- | --- | --- |
| still | < 0.05 g | Sitting, lying, standing still |
| walking | 0.05 – 0.25 g | Light walking, casual movement |
| active | 0.25 – 0.50 g | Brisk walking, light sport |
| exercise | > 0.50 g | Running, HIIT, vigorous activity |

## 6.3 Stage 3: Signal Flagging

The core detection step. For each of the four signal channels, three flag sources are checked in priority order: (1) Z-score thresholds (personalized, only when baseline_ready), (2) Absolute thresholds (universal, always active), (3) Clinical flags from preprocessing.

Z-Score Thresholds (Personalized):

| Signal | Threshold | Direction | Meaning |
| --- | --- | --- | --- |
| HR | z > +2.0 | Elevated | HR is 2+ std devs above personal baseline |
| HRV (RMSSD) | z < -1.5 | Suppressed | HRV dropped 1.5+ std devs below baseline |
| SpO2 | z < -2.0 | Suppressed | Oxygen 2+ std devs below baseline |
| Temperature | z > +2.5 | Elevated | Skin temp 2.5+ std devs above baseline |

Absolute Thresholds (Universal Safety Net):

| Signal | Threshold | Level | Clinical Basis |
| --- | --- | --- | --- |
| HR | >= 120 bpm | High | Resting tachycardia |
| HR | >= 150 bpm | Emergency | Extreme; bypasses persistence |
| SpO2 | <= 94% | Low | BTS/WHO clinical concern |
| SpO2 | <= 90% | Emergency | Hypoxemia; bypasses persistence |
| Skin Temp | >= 37.8 C | High | Possible fever (skin is 4-7C below core) |

Clinical Flag Integration:

| Clinical Flag | Engine Action | Score Assigned |
| --- | --- | --- |
| spo2_hypoxemia (XX%) | Flag SpO2 | 0.85 |
| spo2_clinical_concern (XX%) | Flag SpO2 | 0.55 |
| elevated_skin_temp (XXC) | Flag temp | 0.30 |

## 6.4 Stage 4: Motion-Aware Suppression

If motion is "active" or "exercise", AND only HR is flagged (not SpO2), AND HR < 150 bpm: the HR flag is completely cleared. Elevated HR during exercise is expected.

Critical exception:

If SpO2 is ALSO flagged during exercise, HR suppression does NOT occur. A runner with HR=145 and SpO2=91% is in potential danger — exercise does not explain the oxygen drop.

## 6.5 Stage 5: Raw State Determination

| Flag Count | Raw State | Rationale |
| --- | --- | --- |
| 0 flags | normal | All signals within baseline |
| 1 flag (primary) | stress | Single signal deviation |
| 1 flag (temp only) | normal | Temp alone is too weak — advisory only |
| 2+ flags | distress | Multiple physiological systems disrupted |

## 6.6 Stage 6: Persistence Enforcement

Conditions must persist for 60 seconds before the engine promotes the state. A 2-frame grace period absorbs sensor noise at the boundary.

| Configuration | Value | Effect |
| --- | --- | --- |
| PERSISTENCE_STRESS_S | 60 seconds | Stress conditions must hold 60s before alerting |
| PERSISTENCE_DISTRESS_S | 60 seconds | Distress conditions must hold 60s before escalating |
| grace_frames | 2 frames | Brief signal dropouts do not reset the timer |

Emergency bypass:

SpO2 <= 90% or HR >= 150 at rest bypass persistence entirely and escalate immediately.

## 6.7 Stage 7: Confidence Scoring

Normal state:

confidence = SQI x 0.6 + window_factor x 0.4

Stress/Distress state:

confidence = signal_score x 0.5 + persist_factor x 0.3 + sqi_factor x 0.2

Signal score is the weighted sum of flagged signals. Persist factor ramps from 0.3 to 1.0 over 120 seconds. SQI factor scales with data quality.

# 7. The Three-State Model

| State | Trigger | Severity | User Action |
| --- | --- | --- | --- |
| Normal | 0 primary signals flagged | Low | No intervention needed |
| Stress | 1 primary signal sustained 60s | Low–Medium | Advisory: breathing, relaxation |
| Distress | 2+ primary signals sustained 60s | High | Urgent: reach out for help |

Why 2+ signals for distress: a single elevated signal could have many benign explanations (coffee for HR, altitude for SpO2). When multiple systems are simultaneously disrupted, the probability of genuine distress is much higher.

# 8. Signal Channels Deep Dive

## 8.1 Heart Rate (HR) — Weight: 0.35

Primary stress indicator. Sympathetic activation directly elevates HR within seconds. Z-score threshold: z > +2.0. Absolute: >= 120 bpm (high), >= 150 bpm (emergency). Exercise suppression: YES — HR flag cleared during exercise unless SpO2 also drops.

## 8.2 Blood Oxygen Saturation (SpO2) — Weight: 0.30

Critical safety signal. Normal: 95-100%. Below 94%: BTS/WHO clinical concern. Below 90%: hypoxemia / medical emergency (bypasses persistence). Exercise suppression: NEVER — SpO2 drop during exercise is always concerning.

## 8.3 Heart Rate Variability (HRV RMSSD) — Weight: 0.25

Autonomic nervous system marker. Suppressed HRV indicates reduced parasympathetic tone. Z-score threshold: z < -1.5 (more sensitive than others because HRV captures chronic stress). No absolute threshold — HRV varies too much between individuals for universal limits. Exercise suppression: NOT suppressed — low HRV during exercise can indicate overtraining.

## 8.4 Skin Temperature — Weight: 0.10 (weakest)

Weak supporting signal. Ring measures SKIN temperature (4-7C below core). Affected by ambient temperature, clothing, post-meal vasodilation. Z-score threshold: z > +2.5 (highest threshold). Absolute: >= 37.8C skin. SPECIAL RULE: Temperature ALONE never triggers stress — supporting signal only.

# 9. Dual-Threshold Architecture

| System | When Active | Advantage | Example |
| --- | --- | --- | --- |
| Z-Score (Personalized) | Only after baseline_ready | Catches deviations from individual normal | Athlete HR 48->80 bpm (z=+8) flagged even though 80 is "normal" for most |
| Absolute (Universal) | Always, including cold-start | Safety net for extreme values | HR >= 150 or SpO2 <= 90% flagged regardless of baseline |

# 10. Clinical Flag Integration

| Concept | Reject Reasons | Clinical Flags |
| --- | --- | --- |
| Purpose | Data quality failure | Health advisory |
| Effect on frame | accepted=False, frame discarded | accepted unchanged, frame still analyzed |
| Examples | finger_off, low_ppg_variance | spo2_clinical_concern, elevated_skin_temp |
| Engine behavior | Frame skipped at gate | Flags integrated into signal flagging |

A person with fever and clean sensor data should still have their data analyzed. Clinical concerns are advisory — they never force frame rejection.

# 11. Motion-Aware Exercise Suppression

| Scenario | HR | SpO2 | Motion | HR Flag | Result |
| --- | --- | --- | --- | --- | --- |
| Normal jog | 140 | 97% | exercise | CLEARED | Normal |
| Sprint | 165 | 97% | exercise | KEPT (>=150) | Stress |
| Jog + altitude sickness | 140 | 91% | exercise | KEPT (SpO2 also flagged) | Distress |
| Resting stress | 120 | 97% | still | KEPT | Stress |

# 12. Persistence Tracker

A single-frame anomaly should not generate an alert. The PersistenceTracker ensures conditions must be sustained for 60 seconds before state promotion.

The tracker maintains independent stress and distress timers with a 2-frame grace period. If a condition drops for 1-2 frames (sensor noise), the timer does NOT reset. Only after 3+ consecutive clear frames does the timer reset.

The tracker fully resets on frame rejection (finger off, low SQI) to prevent phantom persistence from building up during data gaps.

# 13. Emergency Bypass Logic

| Condition | Action | Rationale |
| --- | --- | --- |
| SpO2 <= 90% | Immediate DISTRESS (bypass persistence) | WHO hypoxemia definition — every second matters |
| HR >= 150 at rest | Immediate STRESS minimum (bypass persistence) | Abnormal tachycardia while stationary |

The HR emergency bypass does NOT fire during exercise. HR 160 while sprinting is normal; HR 160 while sitting is not.

# 14. Confidence Scoring Model

| Scenario | Signal Score | Persistence | SQI | Confidence |
| --- | --- | --- | --- | --- |
| HR only, just started | 0.21 | 0.3 | 1.0 | 0.40 |
| HR + SpO2, 90 seconds | 0.49 | 0.83 | 1.0 | 0.69 |
| HR + SpO2, 120+ seconds | 0.59 | 1.0 | 1.0 | 0.80 |
| Emergency SpO2=88% | 0.45 | 0.3 | 1.0 | 0.52 |

# 15. STAR-Principle Alert System

| Category | Severity | Example Message |
| --- | --- | --- |
| NORMAL | Low | Everything looks steady. You're doing great. |
| WARMING UP | Low | We're still learning your baseline. |
| TEMP ELEVATED | Low | Your skin temperature is a bit warmer than usual. |
| STRESS (HR) | Low | Your heart rate is a bit higher than usual. |
| STRESS (SpO2) | Medium | Your oxygen level dipped a little. |
| EXERCISE | Low | Your readings are elevated, but you seem to be moving. |
| DISTRESS | High | We're noticing signs of distress. |
| DISTRESS (SpO2) | High | Your oxygen level is lower than expected. |

Design principles:

No medical jargon ("oxygen level" not "SpO2 saturation"). No alarming language. Actionable suggestions (breathing exercises, reaching out). Calm, empathetic tone. No diagnosis — "signs of distress" not "panic attack". Multiple message variants rotate to avoid repetition.

# 16. Cold-Start Handling

When baseline_ready is False: z-score thresholds are disabled (all z-scores = 0.0), only absolute thresholds are active, the user sees "Getting to know you" message, and confidence is lower. Typical warmup: 30-60 seconds at 1 Hz.

Safety during cold-start: if someone puts on the ring and immediately has SpO2 at 88%, the absolute threshold catches it. The safety net is always on.

# 17. Synthetic Training Dataset

| Metric | Value |
| --- | --- |
| Total frames | 9,926 |
| Scenario instances | 2,121 |
| User profiles | 15 (ages 16-65, M/F, sedentary to athletic, plus special populations) |
| Scenario templates | 37 (7 normal, 4 exercise, 6 stress, 8 distress, 10 edge, 2 transition) |
| Columns per frame | 53 |
| Normal : Stress : Distress : Rejected | 54.4% : 21.4% : 22.0% : 2.1% |

# 18. Validation Results

Overall accuracy: 65.8%

| Class | Precision | Recall | F1 |
| --- | --- | --- | --- |
| Normal | 94.6% | 60.9% | 74.1% |
| Stress | 49.0% | 45.9% | 47.4% |
| Distress | 48.3% | 93.9% | 63.8% |
| Rejected | 100.0% | 100.0% | 100.0% |

Key takeaway: Distress recall is 93.9% — the engine catches almost all dangerous conditions. Normal precision is 94.6% — when it says normal, it's almost always right. The synthetic dataset reveals calibration gaps (exercise thresholds, HR-HRV coupling) that inform threshold tuning.

# 19. Calibration Levers

| Parameter | Current | Increase Effect | Decrease Effect |
| --- | --- | --- | --- |
| ZSCORE_TH["hr"] | 2.0 | Fewer HR detections (more specific) | More HR detections (more sensitive) |
| ZSCORE_TH["spo2"] | -2.0 | More SpO2 detections | Fewer SpO2 detections |
| ABS_TH["hr_high"] | 120 bpm | Fewer absolute HR flags | More absolute HR flags |
| ABS_TH["spo2_low"] | 94% | Fewer SpO2 flags | More SpO2 flags |
| PERSISTENCE_STRESS_S | 60s | Slower to alert (fewer FPs) | Faster to alert |
| MIN_SQI_FOR_DETECTION | 0.5 | More frames skipped | More frames evaluated |
| SIGNAL_WEIGHTS["hr"] | 0.35 | HR has more confidence influence | HR has less influence |
| MOTION_TH["active"] | 0.50 | More classified as active | More classified as exercise |

# 20. File Reference

| File | Purpose |
| --- | --- |
| distress_engine.py | Complete standalone distress detection engine (991 lines) |
| generate_synthetic_dataset.py | Generates 9,926 labelled frames + runs validation |
| synthetic_dataset.csv | Flat CSV with 53 columns — every frame labelled |
| synthetic_scenarios.csv | Scenario-level summaries with mean vitals |
| synthetic_dataset_summary.json | Full metadata, profiles, validation results |
| preprocessing_pipeline.py | Upstream: raw sensor to processed output (separate system) |

How to run:

python distress_engine.py                  # 13-scenario demo

python generate_synthetic_dataset.py       # generate dataset + validate

from distress_engine import DistressEngine

engine = DistressEngine()

result = engine.evaluate(pipeline_output)

print(result["state"], result["alert"]["message"])

— End of Document —